## Slide 1: LYRA on-board PROBA2
EUV irradiance inter-calibration workshop
M. Dominique + LYRA team
October 2011, LASP

## Slide 2: PROBA2 orbit
PROBA2 orbit:
Heliosynchronous
Polar
Dawn-dusk
725 km altitude
Duration of 100 min
Occultation season:
  Visible:October-February
  Maximum duration 20 min per orbit

## Slide 3: LYRA highlights
4
5
mm
mm
1: View-limiting aperture (Ø 8mm)
2: precision aperture (Ø 3mm)
3: filter
4: LEDs (λ 375 and 465 nm)
5: detector (Ø 4mm)

## Slide 4: Diamond detectors

## Slide 5: Details of LYRA channels

## Slide 6: Filter + detector responsivity
Unit 1
Unit 2
Unit 3

## Slide 7: Operation: systematic campaigns
 | Integra-tion time | Units | Cover status | LED status | Pointing | Occur-ence
Range | 10ms  10s | max. two at a time | open / close | off / 375nm / 465nm | 0° (Sun)  3° | 
Nominal | 50 ms | U2 | open | off | Sun | N/A
Back-up A | 50ms | U2+3 | open | off | Sun | 1 / 2weeks
Back-up B | 50ms | U2+1 | open | off | Sun | 1 / 3months
Calibra-tion | 50ms | 1) U2+1 2) U2+3 | close | off (DC) on (LED) | N/A | 1/ 2weeks
Paving | 50ms | - | open | off | From 0° to 3° | occasio-nal

## Slide 8: Operation: occasional campaigns
Scientific campaigns
    Observation of flares: unit 2 and 3
    Occultations in all three units
    Eclipses in all three units
Bake-out => no real effect so far

## Slide 9: Long term evolution
Work still in progress …
Various aspects (to be) investigated:
Degradation due to contaminant layer
Dark current evolution (detector degradation)
Response to LED signal acquisition (detector spectral evolution)
An alternative way to probe the spectral evolution (detector + filter): occultations
Flat-field evolution

## Slide 10: Long term evolution
Work still in progress …
Various aspects (to be) investigated:
Degradation due to contaminant layer
Dark current evolution (detector degradation)
Response to LED signal acquisition (detector spectral evolution)
An alternative way to probe the spectral evolution (detector + filter): occultations
Flat-field evolution

## Slide 11: Degradation of unit 2
Time after first light (days)
Uncalibrated signal (counts/ms)
> 95%
> 95%
~ 90%
~ 25%
Corrected signal (counts/ms)
Time after first light (days)

## Slide 12: Degradation of units 1 and 3
30%
15%
/
/
10%
20%
10%
/

## Slide 13: Long term evolution
Work still in progress …
Various aspects (to be) investigated:
Degradation due to contaminant layer
Dark current evolution (detector degradation)
Response to LED signal acquisition (detector spectral evolution)
An alternative way to probe the spectral evolution (detector + filter): occultations
Flat-field evolution

## Slide 14: Dark current evolution
Variations correlated with temperature evolution
I. Dammasch + M. Snow
A. De Groof
Temperature evolution
Dark current in Lyman alpha

## Slide 15: Long term evolution
Work still in progress …
Various aspects (to be) investigated:
Degradation due to contaminant layer
Dark current evolution (detector degradation)
Response to LED signal acquisition (detector spectral evolution)
An alternative way to probe the spectral evolution (detector + filter): occultations
Flat-field evolution

## Slide 16: LED signal time series
Very little change over the mission.
Bimodal appearance is due to systematic difference between first and second measurements during each observing sequence.
Low detector degradation,
if any
M. Snow

## Slide 17: Long term evolution
Work still in progress …
Various aspects (to be) investigated:
Degradation due to contaminant layer
Dark current evolution (detector degradation)
Response to LED signal acquisition (detector spectral evolution)
An alternative way to probe the spectral evolution (detector + filter): occultations
Flat-field evolution

## Slide 18: Probing the evolution of bandpasses: occultations
~ 900 nm
Unit 3

## Slide 19: Long term evolution
Work still in progress …
Various aspects (to be) investigated:
Degradation due to contaminant layer
Dark current evolution (detector degradation)
Response to LED signal acquisition (detector spectral evolution)
An alternative way to probe the spectral evolution (detector + filter): occultations
Flat-field evolution

## Slide 20: Flat-field evolution
1
5
2
3
4
6
7
3mm
Unit 1
Unit 2
(nominal)
Unit 3
5mm substrate

## Slide 21: Non-solar features in LYRA data
Large Angle Rotations
Flat field
LAR: four times an orbit
SAA affects more Si detectors independently of their bandpass

## Slide 22: Non-solar features in LYRA data
Occultation: from mid-October to mid-February
Auroral perturbation
  Only when Kp > 3
  Only affects Al and Zr channels independently of the detector type
  Does not affects SWAP (though observing in the same wavelength range)
Occultations

## Slide 23: Data products
Data products and quicklook viewer on
http://proba2.sidc.be

## Slide 24
Collaborations